- TEST SPECIFICATION

| Job Number | TFS-EMC202506-060 |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| EUT Name | Hot Plate and Stirrer |  |  |  |  |
| EUT Model/SKU Number | SP88857198-EVAL1 |  |  |  |  |
| EUT Serial Number | M3719810042592886 |  |  |  |  |
| Basic Standard | EN 61000-4-11:2020 & IEC 61000-4-11:2020 |  |  |  |  |
| Product Standard | IEC 61326-1:2020, EN 61326-1:2021 |  |  |  |  |
| Immunity test requirement | ☒ Basic ☐ Industrial ☐ Controlled ☐ Custom |  |  |  |  |
| Test Port | ☒ Power Line |  |  |  |  |
| Test Level | Voltage Dips |  |  |  | Interruption |
|  | 0 % | 0 % |  | 70% |  | 0 % |
|  | 0.5 cycle | 1 cycle |  | 25/30cycle |  | 250/300cycle |
| Number of Voltage Dips/Interruptions | ☒ 3 times ☐ Custom______ |  |  |  |  |
| Time between Dips/Interruptions | ☒ 60 sec ☐ Custom______ |  |  |  |  |
| Phase Angle | 0° to 360° with 45° step |  |  |  |  |
| Test Mode | Mode A |  |  |  |  |
| EUT Configuration | ☒ Tabletop |  |  | ☐ Floor standing |  |
| EUT Input Voltage & Frequency | 230 VAC 50/60Hz |  |  |  |  |
| Ambient Temperature (°C) | 26.6 |  |  |  |  |
| Relative Humidity (%) | 64.7 |  |  |  |  |
| Test Date | 02.07.2025 |  |  |  |  |
| Tested by | Md Abubakar Siddeeq |  |  |  |  |

- EQUIPMENTS USED

| S. No | Equipment Name | Make | Model No. | Serial No. | Cal Due Date |
| --- | --- | --- | --- | --- | --- |
| 1 | Compact NX Generator | AMETEK CTS | compact NX5 bspt-1-300-16 | P2237267200 | 11.03.2026 |
| 2 | Motorized Variac | AMETEK CTS | variac NX1-260-16 | P2231266368 | 12.03.2026 |

- EUT MODIFICATION RECORD

The table below details the modifications made to the EUT during the test program.

| Modification state | Description of Modification still fitted to EUT | Modification fitted by | Date modification fitted |
| --- | --- | --- | --- |
| 0 | Initial state |  |  |

- FUNCTIONAL CHECK:

Functional check is conducted as per SOP reference number: IEC-SOP-502.and observed no deviation.

- MONITORING PARAMETERS:

- RPM & Temperature needs to be monitored in the LCD display

- Temperature: 60 ±5° C

- RPM:1500 ±2%

- TEST OBSERVATION

Voltage Dips:230VAC 50Hz

| Test Level % Ut | Duration (cycles/periods) | Observation |
| --- | --- | --- |
| 0 | 0.5 | B1 |
| 0 | 1 | B2 |
| 70 | 25 | B3 |

Interruptions:

| Test Level % Ut | Duration (cycles/periods) | Observation |
| --- | --- | --- |
| 0 | 250 | B6 |

Voltage Dips:230VAC 60Hz

| Test Level % Ut | Duration (cycles/periods) | Observation |
| --- | --- | --- |
| 0 | 0.5 | B4 |
| 0 | 1 | A |
| 70 | 30 | B5 |

Interruptions:

| Test Level % Ut | Duration (cycles/periods) | Observation |
| --- | --- | --- |
| 0 | 300 | B6 |

NOTE 1:

A: No degradation was observed during the test based on the parameters monitored.

B1: The Rpm speed changed to 1500 -1445 while performing the testing but temp within limits

B2 The Rpm speed changed to 1500 -1386 while performing the testing but temp within limits

B3 The Rpm speed changed to 1500 -1257 while performing the testing but temp within limits

B4: The Rpm speed changed to 1500 -1471 while performing the testing but temp within limits

B5: The Rpm speed changed to 1500 -1341 while performing the testing but temp within limits

B6: EUT Display has turned off and again it has recovered back, and manual interruption required to operate the unit.

- TEST SETUP PICTURES

Figure 1: Photograph of Voltage Dips and short Interruptions test setup

- REFERENCE FOLDER:

Product Test Lab/Product Testing/Projects/2025/TFS-EMC-202506-060/Test Data/VDIPS

- SOFTWARE DETAILS:

| Software Used | Software Version |
| --- | --- |
| iec.control | 10.3.2 |

- RESULT:

| Required Performance Criteria | Voltage Dips |  |  | Interruption |
| --- | --- | --- | --- | --- |
|  | 0 % | 0% | 70% | 0 % |
|  | B | B | C | C |
| Met Performance Criteria (230 VAC 50Hz) | B1 | B2 | B3 | B6 |
| Met Performance Criteria (230 VAC 60Hz) | B4 | A | B5 | B6 |

| Tested By |  |
| --- | --- |
| Name | Md Abubakar Siddeeq |
| Signature |  |
| Date | 02.07.2025 |